## Slide 1: UAMV

## Slide 2: UAMV - Agenda
Principle of operation & integration w/ UAM
Technical system overview
Features & UI vs. previous UAMV
Installation & packing list
Preventive maintenance
Review history of service cases & SW
Suggested practicals (to be done afterward)
Questions/comments/feedback

## Slide 3: UAMV – Principle of Operation
  The UAMV sits after the back bar, it creates negative pressure to draw fresh gas over the vaporizers and delivers positive pressure to the patient
  Old UAMV was bellows-based
    During exhalation, bellows rise to draw in fresh gas
    During inhalation, bellows descend to deliver ventilation
  New UAMV is turbine-based
    During inhalation, turbine draws in fresh gas and delivers ventilation

## Slide 4: UAMV – Integration with UAM
  Old UAMV was in
  parallel with manual bellows
Back bar
Breathing block
Bellows
UAMV
  New UAMV is in
  series before manual bellows
Back bar
Breathing block
Bellows
UAMV

## Slide 5: UAMV – Integration with UAM
Breathing block
UAMV manifold inlet
Flex hose connector
UAMV inlet
UAMV outlet
UAMV manifold outlet
From back bar
To patient
  Note: If the UAMV is removed, a tube must be put in place between the UAMV manifold inlet & outlet to avoid a leak

## Slide 6: UAMV – Integration with UAM
  Back bar  UAMV  Breathing block

## Slide 7: UAMV – Features & UI vs. Old UAMV
  UI kept largely to same to reduce training changes
  Additional Features
    Integrated O2 analysis & solenoid control
    Li-ion batteries
    Modes & parameters (PEEP, loops, R, C, etc.)
    Service utilities

## Slide 8: UAMV – Features & UI vs. Old UAMV
    Sleep mode: Disables knob & touchscreen to avoid turning on during shipping
      To enter – 9 short knob presses, then press & hold
      To exit – Connect mains power
    Hard reset: Power cycles the machine if otherwise unresponsive
      5 short knob presses, then press & hold
      To exit – Connect mains power

## Slide 9: Eng. Settings
Settings -> Eng. Settings -> Service Pin -> 7777
Calibration
Voltages Debug
Error Log
Configuration (FiO2 monitor, buzzer alarm)
Device Stats
Text mode
Knob sequences

## Slide 10: Voltages Debug
Includes multiple pages, click to cycle through

## Slide 11: Installation & Packing List
  Accessories, parts packing
  Installing a UAMV on CVI UAMs
  Installing a UAMV on OES UAMs (retrofits)

## Slide 12: Installation – UAMV Accessories
Part | Qty | P/n | Purpose
Valve drive line adapter | 1 | 1146 | Connect UAMV to breathing block port
Valve drive line (1.5m) | 1 | 1216 | Connect UAMV to breathing block port
Flow sensor | 2 | 1532 | Measure patient airway flow & pressure
  UAMV Accessories – packed with UAMV

## Slide 13: Installation – UAMV Accessories
Part | Qty | P/n | Purpose
Power cord (Type G) | 1 | GRN-PEP-100041 | Connect UAMV to UAM auxiliary AC power socket
UAMV Breathing block port w/ o-ring | 1 | GRN-ASM-100078 | Connect drive valve line adapter to UAM breathing block
Inlet/outlet tube | 2 | GRN-PMP-100039 | Connect UAMV inlet/outlet to UAMV manifold
Ethernet cord | 1 | GRN-PEP-100042 | Connect UAMV to nitrous cut-off solenoid
Ethernet cord clip | 1 | GRN-PMP-100040 | Provide strain relief for cut-off solenoid cord
Fuse 2.5 Amp 250V | 2 | GRN-PEP-100043 | Spares, in case inlet fuses burn
  UAMV Accessories Kit – packed with UAM

## Slide 14: Installation (Canon UAM)
Turn on UAMV
Confirm SW version
Confirm O2 analysis is correct
Zero offsets
Calibrate Patient Flow
Connect:
  UAM inlet/outlet tubes
  Valve drive line
  Solenoid cable
  Flow sensor

## Slide 15: Installation – Retrofit (OES UAM)
Goals
  Replicate UAMV manifold
  Connect backbar to UAMV manifold
  Connect valve drive line to breathing block
Parts
  UAMV Accessories Kit – Provide connections for UAM
  Rear Plate or Hatch – Create openings for UAMV manifold
  Kit A – Replicate UAMV manifold
  Kit B – Connect valve drive line adapter to old breathing block port

## Slide 16: Installation – Retrofit (Rear Hatch & Plate)
Rear Hatch & Plate – Create openings for UAMV manifold
UAMV retrofit rear plate (GRN-MFG-100021)
UAMV retrofit rear hatch (GRN-MFG-100020)

## Slide 17: Installation – Retrofit (Kit A)
Kit A – Replicate UAMV Manifold
Parts
Qty 2 of Breathing circuit elbow connector (GRN-PMP-100163)
Qty 2 of Breathing circuit straight connector (GRN-PMP-100164)
Qty 1 of Breathing circuit (short) (GRN-PMP-100162)
Qty 1 of Tube anti-kink clip (GRN-MFG-100001)
Qty 4 of Cable tie, narrow, 8" long (GRN-PMP-100207)
Procedure
Assemble connectors to make elbows
Attach one elbow to the hose going to the breathing block
Attach breathing circuit between backbar and one elbow
Attach anti-kink clip to avoid bends in the breathing circuit
Place elbows through holes in retrofit plate/hatch
Use cable ties to hold elbows in place
Ensure there are no kinks or twists

## Slide 18: Installation – Retrofit (Kit B)
Kit B – Connect valve drive line to breathing block
Parts
Qty 1 Breathing circuit elbow connector (GRN-PMP-100163)
Qty 1 Reusable straight connector, 22mm ID x 5cm (GRN-PMP-100206)
Procedure
Insert the valve drive line adapter into the straight connector
Attach the elbow connector to the UAMV port in the breathing block
Attach the straight connector to the elbow connector
New-style breathing block (reference)
Old-style breathing block

## Slide 19: Installation – Retrofit (Identifying Parts)
Is the UAM serial number after 2304-008?
  If YES, no retrofit parts are required
Is the UAM serial number between A2302-001 and A2304-008?
  If YES, contact the Product team for additional information
Is the UAM serial number before 2302-001?
  If YES, add UAMV Accessory Kit (GRN-ASM-100074)
  If YES, add UAMV retrofit Kit A (GRN-ASM-100073)
Is the UAM serial number before 21810-001?
  If YES, add UAMV retrofit Kit B (GRN-ASM-100072)
Is the UAM serial number before 21712-001?
  If YES, add UAMV retrofit rear plate (GRN-MFG-100021)
  if NO, add UAMV retrofit rear hatch (GRN-MFG-100020)
Is this the first time any UAMV has been installed on the UAM?
  If YES, confirm that the breathing block has a port for a Vent connector (OES p/n 2150-056)
  If YES, confirm whether a Vent connector is needed
Notes: ADD FIRST CVI RUN

## Slide 20: Technical System Overview
  Functional operation
  Sub-assembly layout & serviceability
  Component review
  Discussion of benefits/drawbacks vs. previous design

## Slide 21: Functional Operation
Pneumatic components
  Turbine generates flow
  Inlet check valve prevents backflow
Sensors
  Pressure sensor (ambient)
  Pressure sensor (outlet)
  Pressure sensor (airway)
  Vent flow sensor
  Pt flow sensor (pressure)
Inlet w/ check valve
Outlet
Filter
O2
sensor
Flow
sensor
Pressure
sensors

## Slide 22: Functional Operation
Inlet
Filter
Reservoir
Turbine
V Flow Sensor
Outlet
Fenton
Balloon
O2
Sensor
Valve
Turbine Seal

## Slide 23: Technical System Overview - Electronics
  Boards
    Main display
    Pneumatic control
    Sensor board
    Power management
    Battery pack
    USB interface
  Connections
    Mains power
    Micro USB
    Solenoid (ethernet)
  Benefits
    All sensors on one board
    USB connection to any board

## Slide 24: Casing
4 screws, removes from rear

## Slide 25: Pneumatic Module
Contains all sensors, pneumatic components

## Slide 26: Battery Module
Contains cells, management board

## Slide 27: Battery Module

## Slide 28: Power Module
Contains transformer, power board

## Slide 29: Display Module
Contains screen, driver

## Slide 30: Benefits & Drawbacks of New Design
  Benefits
    Integrated O2 analysis, consolidates electronics
    Allows for setting PEEP
    Increased serviceability
    Lower cost
  Drawbacks
    Adds minimal resistance to draw-over
    Additional potential for leak
    Less easily understood

## Slide 31: UAMV Preventive Maintenance
User Calibrations
Eng. Calibrations

## Slide 32: UAMV Preventive Maintenance
Consumables
  MOX-3 O2 sensor
  Inlet filter
Required equipment & tools
  Hex wrenches
SW upgrade procedure
Batteries
  Lithium ion cells
  Confirm orientation

## Slide 33: Software History Review
Software should always be updated during PMs, if not more urgent
Version 1.04
Display: changed timing settings to address white screen problem
Fenton valve leak test: added filter to reduce low flow noise, Pass threshold: <= 0.3 LPM
GUI: added hours of use counter to Device Stats Window
Pneumatic module: reduced motor block detection time to 50ms
Real Time Clock: Add a bit mask to remove the leap year flag from month data byte
Version 1.03
Version increase to resolve possible confusion between GUI-programmer (MacOS) and terminal programmer (all OS) experienced on 30/11/2023 during troubleshooting in Madagascar)
Version 1.02
Power off via push & hold knob for 3 seconds
Leak test (pass/fail) in user calibration menu
Changed dialog title from “Inspiratory time settings” to “Inspiratory time”
Version 1.01
Added power on beep as “audible” feedback
Added startup delay to handle white screen problem
Changed flow control to be based on patient flow sensor

## Slide 34: UAMV Cases – Fenton Balloon
  Symptom: UAMV provides increasingly large Vt, eventually resets
  Cause:
    Fenton balloon is providing an incomplete seal
    Incomplete seal causes gas to leak out during inspiration
    UAMV attempts to compensate for leak by increasing flow
    Eventually, turbine exceeds maximum power draw and resets
  Fix:
    All units are tested for Fenton balloon location at the factory
    UAMV software updated to prevent overcompensation and resetting

## Slide 35: UAMV Cases – White Screen
  Symptom: During normal operation, screen goes white
  Cause:
    Display controller does not complete frame on time, LCD panel is out of sync, LCD panel power supply shuts down
  Fix:
    UAMV software updated to prevent issue, starting with v1.04

## Slide 36: UAMV Cases – Pneumatic Module Reset
  Symptom: During normal operation, ventilator resets continuously
  Cause:
    TBD, troubleshooting in progress
  Fix:
    Disconnecting or replacing the pneumatic module resolves the behavior

## Slide 37: Suggested Practicals
Installation procedure (CVI UAM)
Installation procedure (all retrofit combinations)
Follow tubing & wires, identify components
Remove & replace each module
Replace battery cells

## Slide 38: Questions / Comments / Feedback